MUI 121022

FilterCheck

User Manual

7506 Holley Circle

Panama City Beach, FL 32408

December 10th, 2022

Table of Contents

Revision Table	3

Overview:	4

User Safety Tips:	4

Device Operation:	5

Preforming a Calibration Check:	9

Charging the Device	11

Directions:	11

Conclusion:	12

# Revision Table

| Date | Revision Number | Change Description | Job Title | Initials |
| --- | --- | --- | --- | --- |
| 12/10/2022 | 0 | Document Created | Software Engineer | LAD |

# Overview:

This Document’s purpose is to give information to a FilterCheck user on the procedures they must follow to operate the FilterCheck and the process of completing a FilterCheck Calibration Check.

# User Safety Tips:

- To avoid damage to FilterCheck components, make sure the device is powered down before shutting the lid.

- It is recommended to use the device in cleaner environments to maximize the devices service life.

- To avoid potential damage or shock do not take the FilterCheck apart.

- Store the device with a dust plug to protect the internal components from dirt.

- To avoid clogs/blockages clean the metal screen that is inserted into the flow port.

- To avoid damage to the electronics do not intentionally submerge in water.

- To avoid damage to the electronics only charge the device at 30v with a minimum of 2A.

# Device Operation:

- Remove the Dust Plug

- Turn on the Power Switch

- A blinking Blue LED represents the device is initializing (estimated initialization time: 45 seconds)

- A solid Green LED represents the device is ready to test filters

- Verify the system to be calibrated:

  - The flow rate must be within the tolerance of 85 ± 1.19 SLPM (83.81 to 86.19 Standard Liters per Minute)

    - Standard Temperature & Pressure (STP) (760 mmHg & 0° Celsius)

  - The low-pressure readings must be within the tolerance of 0 ± .06 cmH2O (-.06 to .06 Centimeters of Water)

- Insert a filter to begin testing

- A blinking Green LED represents the device is analyzing the filter & is adjusting the flow rate back to the required 85 ± 1.19 SLPM (83.81 to 86.19 SLPM)

- The Green LED will stop blinking and turn back solid to represent flow is stable and the test has been completed

- Press the pause button & record the resulting pressure being displayed

- Press the resume button & remove the filter

- a Blue LED will blink representing the device is zeroing

# Preforming a Calibration Check:

- Complete up to Step: 5 (“Insert a filter to begin testing.”) of the Device Operation instructions

- Insert the included Verification Orifice to begin the Calibration Check

- Complete up to Step: 8 (“Press the pause button & record the resulting pressure being displayed”) of the Device Operation instructions

- Verify the measured pressure to be -2.4 ± 10% cmH2O (-2.16 to -2.64 Centimeters of Water)

- Remove the Verification Orifice & a Blue LED will blink representing the device is zeroing

- You have now completed the Calibration Check

- To test filters, continue from Step: 5 (“Insert a filter to begin testing.”) of the Device Operation instructions

# Charging the Device

- There is an internal Rechargeable Battery inside the FilterCheck.

- The FilterCheck has approximately an 8-hour run-time on a single charge.

- The internal Rechargeable Battery has an over-charge cut-off switch that protects the device from being left on the charger to long.

- Recommended to charge while not in use, so the FilterCheck doesn’t have to be plugged in to operate.

- The FilterCheck passively charges the internal Rechargeable Battery when plugged into a wall outlet.

## Directions:

- Plug the 30v Power Supply into a wall outlet.

- Plug the 30v Power Supply’s barrel connector into the FilterCheck

# Conclusion:

This concludes the FilterCheck’s User Manual. For questions and/or concerns regarding the FilterCheck please contact Ator Labs at: LAD@AtorLabs.com.